## REVISON HISTORY
|  | ELPIS IT SOLUTION PRIVATE LIMITED |  | Department: | Embedded System
|  | QUALITY MANAGEMENT SYSTEM |  | Doc Ref: | EITSPL-QR-HWP-19
| Integrated Test Report |  |  | REVISION: | 0
|  |  |  | ISSUE DATE: | 17/02/2025
VERSION | REVISION DATE (DD/MM/YYYY) | DESCRIPTION OF CHANGE | AUTHOR (Name & Title) |  | Review and Approval (Name & Title)
0 |  | First Release | Jagatheswaran J, Test Engineer |  | Venugopal Krishnamurthy (CTO)

## APP TESTCASES
| ELPIS IT SOLUTIONS PRIVATE LIMITED |  | DEPARTMENT | Embedded System
| QUALITY MANAGEMENT SYSTEM |  | DOC REF: | EITSPL-QR-HWP-40
| Integerated Test Report |  | REVISION | 0
|  |  | ISSUEDATE | 17/02/2025
TEST CASE REPORT
Project Name : BSH-VAS
Hardware Version | V-1.1 |  | Test Summary
App version(APP 1.2.3.0) |  |  |  |  | TEST PASS | 0 | #DIV/0! | BSH DEVICE
Date |  |  |  |  | TEST FAIL | 0 | #DIV/0!
Report Prepared By | Jagatheswaran J |  |  |  | TEST BLOCK | 0 | #DIV/0!
Module | TC ID | Scenario | Description | Steps | Expected Result | Verification | Priority | Remarks | Test Status (Pass / Fail / Block)
Login Page UI Validation | TC_UI_001 | Page load | Verify login page loads | Open application | Login screen displayed | UI visible | High
| TC_UI_002 | Logo display | Verify logo | Observe left panel | Logo visible | Correct image | Medium
| TC_UI_003 | Title text | Verify title | Observe | "Access Dashboard" shown | Correct text | High
| TC_UI_004 | Layout alignment | Check UI alignment | Observe page | Proper spacing | No overlap | Medium
| TC_UI_005 | Responsive UI | Resize window | Adjust screen | Layout adapts | No distortion | Medium
Username Field | TC_UN_001 | Empty username | Leave blank | Click Login | Error message | Validation works | High
| TC_UN_002 | Valid username | Enter valid input | Login | Accepted | Input works | High
| TC_UN_003 | Special characters | Enter @#$% | Login | Accept/reject per rule | Validation | Medium
| TC_UN_004 | Max length | Enter long username | Input | Restricted | No overflow | Medium
Password Field | TC_PW_001 | Empty password | Leave blank | Login | Error message | Validation works | High
| TC_PW_002 | Password masking | Enter password | Observe | Masked (****) | Secure display | High
| TC_PW_003 | Toggle visibility | Click eye icon | Toggle | Show/hide password | Works correctly | High
| TC_PW_004 | Valid password | Enter correct password | Login | Accepted | Auth success | High
| TC_PW_005 | Wrong password | Enter wrong password | Login | Error message | Auth fails | High
| TC_PW_006 | Max length(50 character) | Enter long password | Input | Accept/restrict | No crash | Medium
Remember Me Checkbox | TC_RM_001 | Checkbox default | Verify default state | Open page | Unchecked | Correct default | Medium
| TC_RM_002 | Enable checkbox | Click checkbox | Select | Checked | UI updated | Medium
| TC_RM_003 | Persistence | Login with checkbox | Restart app | Username remembered | Stored data | High
| TC_RM_004 | Disable checkbox | Uncheck | Restart | No data stored | Works correctly | Medium
Login Button | TC_BTN_001 | Button click | Click login | With valid input | Login triggered | Navigation happens | High
| TC_BTN_002 | Disabled state | Empty fields | Click login | Error | Validation works | High
| TC_BTN_003 | Multiple clicks | Click repeatedly | Rapid click | Single request sent | No duplicate login | Medium
| TC_BTN_004 | Button UI | Verify style | Observe | Proper color/design | UI consistency | Low
Authentication Logic | TC_AUTH_001 | Valid login | Correct credentials | Login | Redirect to dashboard | Successful login | High
| TC_AUTH_002 | Invalid login | Wrong credentials | Login | Error message | Access denied | High
| TC_AUTH_003 | Empty credentials | No input | Login | Validation error | No login | High
| TC_AUTH_005 | Logout flow | Logout | Reopen | Login required again | Session cleared | Medium
| TC_EDGE_001 | Empty inputs | Blank fields | enter blank fields | Error shown |  | High
| TC_EDGE_002 | Long input | Very long strings | Enter long string | Handled safely |  | Medium
| TC_PERF_001 | Application load | Check page load | Open login page | Loads <2 sec |  | Medium
| TC_PERF_002 | Login response | Check login response | Click login | Response <1 sec |  | High
| TC_PERF_003 | Multiple users | Check login users | Concurrent login | Stable system |  | Medium
| TC_036 | Minimum password length(6 character) | Check password length | Enter short password | Reject |  | High
| TC_037 | Maximum password length(50 char) | Check password length | Enter long password | Accept/restrict |  | Medium
| TC_038 | Whitespace input | Check spaces input | Enter spaces | reject |  | Medium
| TC_039 | Rapid clicks | Check times clicks | Click Save multiple times | Prevent duplicate submission |  | Medium
Add User | TC_AU_001 | Add valid user | Enter username, role, password, confirm → Save | user1 / User / Pass@123 | User added successfully |  | High
| TC_AU_002 | Empty username | Leave username blank | - | Error message |  | High
| TC_AU_003 | Empty password | Leave password blank | - | Error message |  | High
| TC_AU_004 | Password mismatch | Enter different passwords | Pass@123 / Pass@124 | Error shown |  | High
| TC_AU_005 | Duplicate username | Enter existing "admin" | admin | Duplicate error |  | High
| TC_AU_006 | Role dropdown validation | Don’t select role | - | Show validation |  | Medium
| TC_AU_007 | Rapid multiple click onsAdmin page | Click on admin page | click multiple time | only one page is open |  | Medium |  |  | previous version When i clicks the Admin Page button multiple times rapidly, the application opens a new Admin Page tab for every click.in this vixed.ersion it is
| TC_AU_009 | User delete | After adding a user in User Management, the user can be deleted. | click delete | user deleted |  | Medium
| TC_AU_010 | Max length username | Enter >50 chars | - | Restrict input |  | Medium
| TC_AU_011 | Save button click | Click Save | Valid data | Action triggered once |  | High
| TC_AU_012 | Rapid multiple clicks | Click Save multiple times | - | Only one user added |  | High
| TC_AU_013 | Field reset after save | Add user | - | Fields cleared |  | Medium
User Management | TC_UM_001 | Table load | Open Admin Page | Table visible |  |  | High
| TC_UM_002 | Default admin user | Check row | "admin" displayed |  |  | High
| TC_UM_003 | Role correctness | Verify role column | Correct role shown |  |  | High
| TC_UM_004 | Add user reflection | Add new user | Table updates instantly |  |  | High
| TC_UM_005 | Column headers | Check headers | Username / Role / Actions |  |  | Medium
| TC_UM_006 | Table alignment | Inspect UI | Proper alignment |  |  | Low
| TC_AC_001 | Edit user | Click edit | Edit popup opens |  |  | Medium
| TC_AC_002 | Delete user | Click delete | Confirmation shown |  |  | High
| TC_AC_003 | Delete admin restriction | Try delete admin | Block action |  |  | High
| TC_AC_004 | Action button visibility | Check UI | Icons visible |  |  | Medium
Page control | TC_PC_001 | Close button (X) | Click top-right X | Page closes |  |  | High
| TC_PC_002 | Navigation tabs | Click Waterfall/Graph | Navigate correctly |  |  | Medium
| TC_PC_003 | Page title | Check title | "Vibration Analyser" shown |  |  | Low
| TC_SEC_004 | Unauthorized access | Non-admin login | Access denied |  |  | High
| TC_UI_001 | Button color consistency |  | Purple buttons uniform |  |  | Low
| TC_UI_002 | Input field alignment |  | Proper spacing |  |  | Low
| TC_UI_003 | Eye icon placement |  | Correct alignment |  |  | Low
| TC_UI_004 | Responsive UI |  | Works in resize |  |  | Medium
| TC_UI_005 | Error message visibility |  | Clear and readable |  |  | High
Popup & UI Validation | TC_UI_001 | Popup open | Verify popup opens | Click device config | Popup visible | UI displayed | High
| TC_UI_002 | Popup close (X) | Close using X | Click X | Popup closes | Disappears | High
| TC_UI_003 | Layout alignment | Check UI alignment | Observe fields | Proper spacing | No overlap | Medium
| TC_UI_004 | Labels visibility | Verify labels | Observe UI | All labels readable | Correct text | High
| TC_UI_005 | Default values | Check default data | Open popup | Fields pre-filled | Matches device | High
Channel Selection | TC_CH_001 | Channel display | Verify channel name | Open popup | "Channel 4" shown | Matches selection | High
| TC_CH_002 | Channel change (if editable) | Select different channel | Change dropdown | Channel updates | UI reflects | Medium
| TC_CH_003 | Channel consistency | Compare with device panel | Observe | Same channel config | Data consistent | High
Enable / Disable Toggle | TC_EN_001 | Default state | Check toggle | Open popup | ON selected | Correct state | High
| TC_EN_002 | Toggle OFF | Click OFF | Disable channel | OFF active | UI updates | High
| TC_EN_003 | Toggle ON | Click ON | Enable channel | ON active | UI updates | High
| TC_EN_004 | Graph effect | Disable channel | Observe graph | Channel hidden | Graph updated | High
| TC_EN_005 | Persistence | Save and reopen | Check state | Same state retained | Stored correctly | High
Name Field | TC_NAME_001 | Default name | Verify name | Open popup | "Channel 4" | Matches channel | High
| TC_NAME_002 | Edit name | Change name | Enter new name | Value updated | UI reflects | High
| TC_NAME_003 | Empty name | Clear field | Save | Error message | Validation works | High
| TC_NAME_004 | Special characters | Enter @#$% | Save | Accept/reject based on rule | Validation | Medium
| TC_NAME_005 | Max length | Enter long text | Save | Restrict length | No overflow | Medium
Color Selection | TC_CLR_001 | Default color | Check color | Open popup | Purple shown | Matches graph | High
| TC_CLR_002 | Change color | Select new color | Pick color | Color updates | UI reflects | High
| TC_CLR_003 | Graph sync | Change color | Save | Graph updates color | Match graph | High
Measurement Mode | TC_MEAS_001 | Default mode | Verify mode | Open popup | "Acceleration" | Correct value | High
| TC_MEAS_002 | Dropdown open | Click dropdown | Open list | Options visible | UI works | High
| TC_MEAS_003 | Change mode | Select new mode | Choose option | Updated | UI reflects | High
| TC_MEAS_004 | Invalid mode | Force invalid | Enter invalid | Reject | Validation works | Medium
Unit Selection | TC_UNIT_001 | Default unit | Verify unit | Open popup | "g" | Correct unit | High
| TC_UNIT_002 | Change unit | Select unit | Choose option | Updated | UI reflects | High
| TC_UNIT_003 | Unit consistency | Check graph | Observe | Matches measurement | Consistent | High
Min / Max Range | TC_RANGE_001 | Default values | Check min/max | Open popup | 0 / 100 | Correct values | High
| TC_RANGE_002 | Valid range | Enter valid values | Save | Accepted | Stored | High
| TC_RANGE_003 | Min > Max | Enter 100 / 0 | Save | Error shown | Validation | High
| TC_RANGE_004 | Negative values | Enter -10 | Save | Reject or handle | Validation | High
| TC_RANGE_005 | Non-numeric input | Enter "abc" | Save | Reject | Validation | High
| TC_RANGE_006 | Boundary values(0-100) | Enter limits | Save | Accepted | Works | Medium |  |  | Taking more than 100 numbers
Save & Cancel Buttons | TC_BTN_001 | Save changes | Click Save | Save config | Data stored | Reopen verify | High
| TC_BTN_002 | Cancel changes | Click Cancel | Close popup | Changes not saved | Reopen check | High
| TC_BTN_003 | Save validation | Invalid input | Click Save | Error message | Validation works | High
| TC_BTN_004 | Multiple clicks | Click Save rapidly | Prevent duplicates | Stable | Medium
IP Address Configuration | TC_IP_001 | Default IP | Verify IP fields | Open popup | 192.168.1.90 shown | Matches device | High
| TC_IP_002 | Valid IP input | Enter valid IP | 192.168.0.10 | Accepted | Saved correctly | High
| TC_IP_003 | Invalid IP (>255) | Enter 300.168.1.1 | Save | Error shown | Validation works | High
| TC_IP_004 | Negative IP | Enter -1 | Save | Reject | Validation works | High
| TC_IP_005 | Non-numeric IP | Enter "abc" | Save | Reject | Validation works | High
| TC_IP_006 | Partial IP | Leave one field empty | Save | Error shown | Validation works | High
| TC_IP_007 | Boundary values | Enter 0 / 255 | Save | Accepted | Valid IP range | Medium
Device Name | TC_NAME_001 | Default name | Verify name | Open popup | "Device 1" | Matches device | High
| TC_NAME_002 | Edit name | Change name | Enter "Device A" | Updated | UI reflects | High
| TC_NAME_003 | Empty name | Clear field | Save | Error shown | Validation works | High
| TC_NAME_004 | Special characters | Enter @#$% | Save | Accept/reject | Validation | Medium
| TC_NAME_005 | Max length | Enter long string | Save | Restrict input | No overflow | Medium
Port Configuration | TC_PORT_001 | Default port | Verify value | Open popup | 5008 | Correct | High
| TC_PORT_002 | Valid port | Enter 8080 | Save | Accepted | Stored | High
| TC_PORT_003 | Invalid port (>65535) | Enter 70000 | Save | Error | Validation works | High
| TC_PORT_004 | Negative port | Enter -1 | Save | Reject | Validation works | High
| TC_PORT_005 | Non-numeric | Enter "abc" | Save | Reject | Validation works | High
| TC_PORT_006 | Boundary ports | Enter 0 / 65535 | Save | Accept valid | Works correctly | Medium
Server IP Configuration | TC_SIP_001 | Default server IP | Verify value | Open popup | 192.168.1.137 | Correct | High
| TC_SIP_002 | Valid server IP | Enter valid IP | Save | Accepted | Stored | High
| TC_SIP_003 | Invalid server IP | Enter 999.x.x.x | Save | Error | Validation works | High
| TC_SIP_004 | Empty server IP | Leave blank | Save | Error | Validation works | High
Sampling Frequency | TC_SF_001 | Default value | Verify frequency | Open popup | 22000 Hz | Correct | High
| TC_SF_002 | Valid input | Enter 800000max | Save | Accepted | Stored | High
| TC_SF_003 | Negative value | Enter -100min | Save | not accepting | Validation | High
| TC_SF_004 | Non-numeric | Enter abc | Save | Reject | Validation | High
Acquisition Interval | TC_AI_001 | Default interval | Verify value | Open popup | 1000 ms | Correct | High
| TC_AI_002 | Valid interval | Enter 50 | Save | Accepted | Stored | High
| TC_AI_003 | Zero interval | Enter 0 | Save | Reject | Validation | High
Trigger Type | TC_TR_001 | Default selection | Verify default | Open popup | GUI Driven selected | Correct | High |  |  | On the first time open the app default trigger type not selecting.Without selecting the trigger type, the configure command is being sent.
| TC_TR_002 | Change trigger | Select Pulse Count | Click | Selected | UI updates | High
| TC_TR_003 | Mutual exclusivity | Select multiple | Try | Only one active | Correct behavior | High
Save & Cancel Buttons | TC_BTN_001 | Save valid data | Click Save | Valid inputs | Data saved | Reopen verify | High
| TC_BTN_002 | Cancel action | Click Cancel | Close popup | No changes saved | Verify revert | High
| TC_BTN_003 | Save invalid data | Click Save | Invalid input | Error shown | Validation works | High
| TC_BTN_004 | Multiple clicks | Click Save repeatedly | Prevent duplicate | Stable |  | Medium
| TC_BTN_005 | Empty all fields | Blank inputs |  | Reject |  | High
Waterfall Graph View | TC_WF_001 | Graph load | Verify graph loads | Open Waterfall Graph tab | Graph area visible | UI displayed | High
| TC_WF_002 | Empty state | No data loaded | Open page | Blank graph / placeholder | No crash | High
| TC_WF_003 | Data rendering | Load data | Click LoadData | Waterfall plot appears | Visual verification | High
| TC_WF_004 | Axis labels | Check axes | Observe graph | Proper labels (freq/time/amplitude) | Matches design | High
| TC_WF_005 | Color gradient | Verify color mapping | Observe graph | Gradient represents amplitude | Matches legend | High
| TC_WF_006 | Real-time update | Start acquisition | Click Start | Graph updates continuously | Smooth refresh | High
| TC_WF_007 | High data load | Load large dataset | Observe | No lag or crash | Performance stable | Medium
| TC_WF_008 | No signal input | Zero data | Load data | Flat/empty graph | Correct behavior | Medium
View Control Buttons (Top / Side / Front / Change Color) | TC_VIEW_001 | Top view | Change to top | Click "top" | Graph view changes | Orientation updated | High
| TC_VIEW_002 | Side view | Change to side | Click "side" | View updated | Different perspective | High
| TC_VIEW_003 | Front view | Change to front | Click "front" | View updated | Perspective changes | High
| TC_VIEW_004 | Change color | Modify color scale | Click "Change Color" | Color palette changes | Graph updated | High
| TC_VIEW_005 | Multiple toggles | Switch views rapidly | Click repeatedly | No crash | Stable UI | Medium
Graph vivew | TC_DV_001 | Device display | Verify device list | Open panel | Devices visible | Correct entries | High
| TC_TS_002 | Multi-channel plotting | Check 8 channels | Load data | 8colored signals shown | Match legend | High
| TC_TS_003 | Axis labels | Validate axes | Observe graph | X=Milliseconds, Y=Amplitude | Correct text | High
| TC_TS_004 | Data accuracy | Validate waveform | Compare CSV | Matches input | Cross-check | High |  |  | some times I observed wrong timeseries plotted .And distorted wave form
| TC_TS_005 | RMS/Peak table | Verify stats | Observe panel | Correct RMS/Peak values | Compare manually | High |  |  | In the timeseries RMS, Peak, and Peak-to-Peak values box is blocking the graph when all channels are selected
| TC_TS_006 | Zoom feature | Click expand icon | Zoom graph | Enlarged view | UI response | Medium
| TC_TS_007 | Signal variation | Load dynamic data | Observe | Smooth waveform changes | No freeze | High
| TC_TS_008 | Channel color consistency | Compare colors | Observe graph | Matches legend | Visual match | High
| TC_TS_009 | Empty data handling | Load empty file | Open graph | No crash, show blank | Stable | High
Spectrogram Graph | TC_SP_001 | Graph rendering | Verify spectrogram loads | Open page | Graph visible | Visual check | High
| TC_SP_002 | Frequency axis | Validate X-axis | Observe | Frequency scale correct | Units shown | High |  |  | some times I observed wrong frequency value I need to verify this on final release
| TC_SP_003 | FFT peaks | Validate frequency peaks | Load known signal | Peaks at expected freq | Compare FFT | High
| TC_SP_004 | Channel legend | Verify legend display | Observe | Channel 1–4 shown | Matches colors | High
| TC_SP_005 | No signal case | Load zero data | Observe | Flat/empty graph | No crash | Medium
| TC_SP_006 | Large frequency data | Load high freq | Observe | Proper scaling | No distortion | Medium
Device View Panel | TC_DV_002 | Active status | Check red indicator | Observe | Shows device status | Matches connection | High
| TC_DV_003 | Device selection | Click device | Select row | Highlighted | UI response | High
| TC_DV_004 | IP validation | Verify IP format | Observe | Valid IP displayed | Format correct | High
| TC_DV_005 | Expand/collapse device | Click arrow | Toggle channels | Expand/collapse works | UI correct | High
Channel Configuration | TC_CH_001 | Channel list | Verify channels | Expand device | Channels 1–4 visible | Correct count | High
| TC_CH_002 | Enable toggle | Turn ON/OFF | Click toggle | State changes | UI reflects | High
| TC_CH_003 | Disable effect | Turn OFF channel | Observe graph | Channel removed | Graph updates | High
| TC_CH_004 | Color mapping | Verify color box | Observe | Matches graph | Visual match | High
| TC_CH_005 | Measurement type | Check "acceleration" | Observe | Correct label | Matches config | High
| TC_CH_006 | Min/Max values | Validate limits | Observe | Correct values | Matches config | High
Acquisition Settings | TC_ACQ_001 | Sampling frequency | Verify value | Observe | 22000 Hz | Correct | High
| TC_ACQ_002 | Acquisition interval | Verify value | Observe | 100-1000 ms | Correct | High
| TC_ACQ_003 | Acquisition duration | Verify value | Observe | 1-10 sec | Correct | High
| TC_ACQ_004 | Invalid input | Enter invalid | Save | Error message | Validation works | High
Trigger Type | TC_TR_001 | Default trigger | Check selection | Open panel | GUI Driven selected | Correct | High
| TC_TR_002 | Change trigger | Select Pulse Count | Click | Updated | Only one selected | High
| TC_TR_003 | Mutual exclusivity | Try multiple | Select all | Only one active | Correct behavior | High
Control Buttons | TC_BTN_001 | Start button | Start acquisition | Click Start | Data starts | Graph updates | High
| TC_BTN_002 | Load Data | Load file | Click LoadData | Graph displays data | Correct rendering | High
| TC_BTN_003 | Button validation | Click without config | Click Start | Error or disabled | Validation works | Medium
| TC_BTN_004 | Rapid clicks | Click multiple times | Observe | No duplicate actions | Stable | Medium
Status Bar | TC_ST_001 | Connection status | Check bottom | Observe | "UDP receiver started..." | Correct state | High
| TC_ST_002 | Data source | Verify file path | Load data | Path shown | Matches file | High
| TC_ST_003 | Time display | Check time | Observe | Updates correctly | Accurate | Medium
Performance Test Cases | TC_PERF_001 | Real-time streaming | Start acquisition | Smooth graph | No lag
| TC_PERF_002 | Multiple devices | Load devices | Stable UI | Medium
| TC_PERF_003 | View switching | Change views | Instant response | Medium
| TC_PERF_004 | No device connected | No devices | Open panel | Show message |  |  |  |  | If the both device is not connected only device status showing red if one device disconnected the app not showing. Red it is still in green.
| TC_PERF_005 | Disable all channels | All OFF | Observe graph | Blank graph
| TC_PERF_006 | Large dataset | Big file | Load data | No lag
| TC_PERF_007 | Corrupted file | Invalid CSV | Error message | High
Alert | TC_AL_001 | Alert display | Verify alert panel | Open alert tab | Panel visible | UI check | High
| TC_AL_002 | Channel name | Verify channel | Observe | Correct channel shown | Matches graph | High
| TC_AL_003 | Frequency threshold | Check threshold | Observe | Value displayed | Matches config | High
| TC_AL_004 | Observed value | Verify value | Compare with graph | Correct value | Data match | High
| TC_AL_005 | Timestamp | Verify time | Observe | Correct timestamp | Matches system time | Medium
| TC_AL_006 | RMS alert | Verify RMS value | Observe | Correct RMS displayed | Match table | High
| TC_AL_007 | Range validation | Check range | Observe | Correct range shown | Matches config | High
| TC_AL_008 | Alert trigger | Exceed threshold | Input high value | Alert generated | Visible in panel | High
| TC_AL_009 | Multiple alerts | Trigger multiple | Load varied data | All alerts listed | Proper logging | Medium
Data Source & Status Bar | TC_DS_001 | File path display | Verify path | Load CSV | Path shown correctly | Matches file | High |  |  | After changing the folder path, closing and reopening the app resets it to the default path.it is fixed in this version
| TC_DS_002 | Connection status | Check status | Observe bottom | "Not Connected" shown | Correct state | High
| TC_DS_003 | Timestamp update | Observe clock | Wait | Time updates | Real-time | Medium
UI Controls | TC_UI_001 | Close button | Click X | Close panel | Window closes | Works correctly | High
| TC_UI_002 | Tab navigation | Switch tabs | Click Graph/Alert | Switch works | Correct page | High
| TC_UI_003 | Layout alignment | Check UI | Observe | Proper alignment | No overlap | Medium
Performance Test | TC_PERF_001 | Load large dataset | Load big CSV | No lag/crash | High
| TC_PERF_002 | Real-time streaming | Continuous input | Smooth graph update | High
| TC_PERF_003 | Multiple channels | Load all channels | No delay | Medium
| TC_EDGE_001 | Corrupted file | Invalid CSV | Error message | High
| TC_EDGE_002 | Missing values | Partial data | Handle gracefully | Medium
| TC_EDGE_003 | Extreme values | Very high amplitude | No overflow | High
| TC_EDGE_004 | Zero data | All zeros | Flat graph | Medium
Threshold View Module Spectrogram | TC_CH_001 | Default selection | Verify default channel | Open page | Channel 1 selected | Dropdown value | High
| TC_CH_002 | Change channel | Select another channel | Choose Channel 2 | Fields update accordingly | UI reflects | High
| TC_CH_003 | Channel persistence | Save and reopen | Select channel | Last selection retained | Stored value | Medium
| TC_CH_004 | Invalid selection | Force invalid input | Try invalid | Reject | Validation | Medium
Configuration | TC_SP_001 | Input center frequency | Enter valid value | 1000 Hz | Accepted | Field updated | High
| TC_SP_002 | Range min/max valid | Enter min < max | 900 / 1100 | Accepted | Correct storage | High
| TC_SP_003 | Range min > max | Enter invalid | 1200 / 1000 | Error shown | Validation | High
| TC_SP_004 | Amplitude input | Enter value | 500 | Accepted | Stored | High
| TC_SP_005 | Color selection | Select color | Choose color | Updated | Matches UI | High
| TC_SP_006 | Empty fields | Leave blank | Click Save | Error message | Validation | High
| TC_SP_007 | Non-numeric input | Enter "abc" | Save | Reject | Validation | High
| TC_SP_008 | Boundary values | Enter 0 / max | Save | Accepted if valid | Works | Medium
Time Series Threshold Configuration | TC_TS_001 | RMS range input | Enter min/max | 7500 / 10000 | Accepted | Stored | High
| TC_TS_002 | Peak range input | Enter min/max | 150 / 200 | Accepted | Stored | High
| TC_TS_003 | Peak-to-peak input | Enter values | 200 / 250 | Accepted | Stored | High
| TC_TS_004 | Invalid range | Min > Max | Enter wrong | Error message | Validation | High
| TC_TS_005 | Empty fields | Leave blank | Save | Error | Validation | High
| TC_TS_006 | Non-numeric input | Enter text | Save | Reject | Validation | High
| TC_TBL_001 | Table display | Verify table loads | Open page | Table visible | UI check | High
| TC_TBL_002 | Data correctness | Verify values | Compare inputs | Matches entered data | Cross-check | High
| TC_TBL_003 | Multiple entries | Add multiple rows | Save data | All rows displayed | Correct order | High
| TC_TBL_004 | Color mapping | Verify color column | Observe | Matches selected color | Visual match | High
| TC_TBL_005 | Table display | Verify table loads | Open page | Table visible | UI check | High
| TC_TBL_006 | Data accuracy | Compare values | Observe | Matches inputs | Correct values | High
| TC_TBL_007 | Update entry | Modify values | Save | Table updated | Reflect changes | High
Action Buttons (Save / Clear) | TC_BTN_001 | Save spectrogram | Click Save | Valid input | Data saved | Table updated | High
| TC_BTN_002 | Save time series | Click Save | Valid input | Data saved | Table updated | High
| TC_BTN_003 | Clear button | Click Clear | Fields filled | Fields cleared | Empty inputs | High
| TC_BTN_004 | Save invalid input | Click Save | Invalid data | Error shown | Validation works | High
| TC_BTN_005 | Multiple clicks | Rapid click | Click repeatedly | No duplicate entries | Stable | Medium
Trigger functionality | TRG_001 | Trigger on Device 1 only | Connect Device 1 & Device 2 → Configure trigger for Device 1 → Start acquisition | Device 1 should trigger and acquire data successfully | Verify Device 1 waveform/data capture starts; Device 2 remains unchanged |  | High
| TRG_002 | Trigger on Device 2 only | Connect both devices → Configure trigger for Device 2 → Start acquisition | Device 2 should trigger successfully | Verify Device 2 data capture starts; Device 1 unaffected |  | High
| TRG_003 | Trigger both devices simultaneously | Connect both devices → Configure trigger for both → Start acquisition | Both devices should trigger independently | Verify both devices capture data correctly |  | High
| TRG_004 | Cross-trigger isolation | Configure trigger on Device 1 only → Start | Device 2 should not trigger automatically | Verify Device 2 remains idle |  | High
| TRG_005 | Stop trigger on Device 1 only | Start trigger on both devices → Stop Device 1 trigger | Device 1 trigger stops; Device 2 continues | Verify Device 2 still captures data |  | High
| TRG_006 | Stop trigger on Device 2 only | Start both → Stop Device 2 trigger | Device 2 stops; Device 1 continues | Verify Device 1 unaffected |  | High
| TRG_007 | Invalid trigger configuration | Enter invalid trigger values → Start | Validation/error message displayed | Confirm trigger does not start |  | High
| TRG_008 | Trigger without device connection | Configure trigger → Disconnect devices / no connection → Start | Error message shown | Confirm no trigger operation starts |  | High
| TRG_009 | Device 1 disconnect during trigger | Start triggered acquisition → Disconnect Device 1 | Error shown for Device 1; Device 2 unaffected | Verify Device 2 continues |  | High
| TRG_010 | Device 2 disconnect during trigger | Start triggered acquisition → Disconnect Device 2 | Device 1 continues normally | Verify Device 1 unaffected |  | High
| TRG_011 | Trigger signal only on Device 1 | Apply trigger signal to Device 1 only | Only Device 1 should trigger | Verify Device 2 does not respond |  | High
| TRG_012 | Trigger signal only on Device 2 | Apply trigger signal to Device 2 only | Only Device 2 should trigger | Verify Device 1 unaffected |  | High
| TRG_013 | Different trigger conditions on both devices | Configure different trigger thresholds/types | Each device should trigger based on its own settings | Compare actual trigger behavior with config |  | High
| TRG_014 | Simultaneous different signals | Apply different signals to both devices | Correct trigger response per device | Verify independent triggering |  | Medium
| TRG_015 | Rapid trigger start/stop cycles | Repeatedly start/stop trigger on both devices | No crash/disconnect | Verify app stability |  | High
| TRG_016 | Trigger during long-duration acquisition | Run triggered acquisition for hours | Stable operation without logout/crash | Monitor app/device stability |  | Medium
| TRG_017 | App background/foreground during trigger | Start trigger → Minimize app → Reopen | Trigger state behaves correctly | Verify acquisition continuity |  | High
| TRG_018 | Trigger timeout/no signal condition | Configure trigger but do not provide signal | App waits correctly or shows timeout behavior | Verify no false trigger |  | High
| TRG_019 | Wrong waveform routing during trigger | Apply signal to Device 1 only | Device 2 should not show Device 1 waveform | Verify waveform isolation |  | High
| TRG_020 | Report/timestamp separation | Trigger acquisitions on both devices → Generate reports | Correct device-specific timestamps/data | Verify report data accuracy
| NTRG_SW_001 | Trigger without device connection | Press trigger switch without connecting devices | Error message / no trigger action | Confirm no acquisition starts |  | High
| NTRG_SW_002 | Trigger switch pressed before Start | Connect devices → Press trigger switch before starting acquisition | Trigger should not activate acquisition unexpectedly | Confirm no false trigger |  | High
| NTRG_SW_003 | Multiple rapid switch presses | Press trigger switch rapidly multiple times | Only one valid trigger event should be detected | Confirm no duplicate triggers |  | High
| NTRG_SW_004 | Switch bounce/noise trigger | Press switch partially / simulate bouncing | System should debounce correctly | Confirm no false multiple triggers |  | High
| NTRG_SW_005 | Long press on trigger switch | Press and hold trigger switch for several seconds | Only one trigger event should occur (as per design) | Confirm no repeated false triggers |  | High
| NTRG_SW_006 | Trigger without pulse signal | Start acquisition without pressing switch | Trigger should wait for valid pulse | Confirm no false trigger |  | High
| NTRG_SW_007 | Trigger signal on Device 1 affects Device 2 | Press Device 1 trigger switch only | Device 2 should not trigger | Confirm trigger isolation |  | High
| NTRG_SW_008 | Trigger signal on Device 2 affects Device 1 | Press Device 2 trigger switch only | Device 1 should remain idle | Confirm isolation |  | High
| NTRG_SW_009 | Trigger during device disconnect | Start trigger mode → Disconnect device → Press switch | Proper error handling | Confirm no crash |  | High
| NTRG_SW_010 | Trigger during app background state | Start trigger → Minimize app → Press switch → Reopen | App should behave as expected | Verify trigger state |  | Medium
| NTRG_SW_011 | Trigger after Stop | Start → Stop → Press trigger switch | No acquisition should start unexpectedly | Confirm trigger ignored |  | High
| NTRG_SW_012 | Trigger while acquisition already running | Start acquisition → Press switch repeatedly | No duplicate/abnormal trigger events | Confirm stable operation |  | High
| NTRG_SW_013 | Broken/stuck switch condition | Simulate continuously ON switch | System should handle gracefully | Confirm no endless triggering |  | High
| NTRG_SW_014 | Delayed switch response | Press switch and monitor response | Trigger should occur within expected response time | Verify timing behavior |  | Medium
| NTRG_SW_015 | Wrong waveform after switch trigger | Press trigger switch | Correct waveform should appear for correct device only | Verify waveform routing |  | High
Start/stop functionality | SS2_001 | Start acquisition on Device 1 | Verify Device 1 starts independently | Connect Device 1 & Device 2 → Tap Start for Device 1 | Device 1 starts acquisition successfully | Verify Device 1 graph/data updates | High
| SS2_002 | Start acquisition on Device 2 | Verify Device 2 starts independently | Connect both devices → Tap Start for Device 2 | Device 2 starts acquisition successfully | Verify Device 2 graph/data updates | High
| SS2_003 | Start both devices simultaneously | Verify simultaneous acquisition | Connect both devices → Tap Start for both | Both devices start independently | Verify both graphs update correctly | High
| SS2_004 | Stop Device 1 only | Verify isolated stop control | Start both devices → Stop Device 1 | Device 1 stops; Device 2 continues | Confirm Device 1 stops, Device 2 active | High
| SS2_005 | Stop Device 2 only | Verify isolated stop control | Start both devices → Stop Device 2 | Device 2 stops; Device 1 continues | Confirm Device 2 stops, Device 1 active | High
| SS2_006 | Stop both devices | Verify stopping both acquisitions | Start both devices → Stop both | Both devices stop correctly | Confirm both graphs stop updating | High
| SS2_007 | Rapid Start/Stop on Device 1 | Verify stability under repeated actions | Perform multiple Start/Stop cycles on Device 1 | No crash/disconnect | Monitor app/device stability | High
| SS2_008 | Rapid Start/Stop on Device 2 | Verify stability under repeated actions | Perform multiple Start/Stop cycles on Device 2 | Stable behavior | Confirm no crash or unexpected behavior | High
| SS2_009 | Start Device 1 while Device 2 running | Verify independent control | Start Device 2 → Start Device 1 | Both work independently | Confirm both acquisitions active | High
| SS2_010 | Stop Device 1 while Device 2 updating | Verify no impact on Device 2 | Start both → Stop Device 1 | Device 2 continues plotting | Confirm Device 2 unaffected | High
| SS2_011 | Device 1 disconnect during acquisition | Verify disconnect isolation | Start both → Disconnect Device 1 | Device 1 error shown; Device 2 unaffected | Confirm Device 2 continues normally | High
| SS2_012 | Device 2 disconnect during acquisition | Verify disconnect isolation | Start both → Disconnect Device 2 | Device 1 continues normally | Confirm Device 1 unaffected | High
| SS2_013 | Device reconnect after disconnect | Verify recovery after reconnect | Disconnect one device → Reconnect → Tap Start | Acquisition resumes normally | Confirm reconnect functionality | Medium |  |  | After opening the app, device connection issues occur intermittently. Sometimes the device connects automatically, and sometimes the app needs to be closed and reopened to establish the connection.
| SS2_014 | Cross-device Start command issue | Verify no unintended cross-trigger | Start Device 1 only | Device 2 should not start | Confirm Device 2 remains idle | High
| SS2_015 | Cross-device Stop command issue | Verify stop isolation | Stop Device 1 only | Device 2 should continue running | Confirm Device 2 remains active | High
| SS2_016 | Wrong waveform routing check | Verify correct waveform mapping | Apply signal to Device 1 only | Device 1 waveform only displayed | Confirm Device 2 shows no Device 1 data | High
| SS2_017 | Device 2 signal isolation | Verify signal isolation | Apply signal to Device 2 only | Device 2 waveform only displayed | Confirm Device 1 unaffected | High
| SS2_018 | Simultaneous signal handling | Verify multi-signal accuracy | Apply different signals to both devices | Correct waveform shown for each device | Compare input vs displayed waveform | High |  |  | We have only one signal generator, so at a time we can provide the signal to only one device.
| SS2_019 | App stability during simultaneous operation | Verify long-duration stability | Run both devices for several hours | No crash/logout/disconnect | Monitor app performance and connection | High |  |  | Over the last 6 days, I observed app crash issues. but, yesterday I kept the device for overnight testing, and the issue did not occur with app version v1.1.0.0.After the one more test I will update this testcase. No abnormal crashes happened.But still not confident in stability
| SS2_020 | Timestamp/data separation | Verify report accuracy | Generate reports for both devices | Correct device-specific timestamps/data | Compare report content with acquisition session | High |  |  | needs to be verify
Configuration | CFG_001 | Verify Configure screen opens successfully | Open MDAQ app → Connect device → Navigate to Configure screen | Configure screen opens without crash | Confirm screen loads properly and all fields are visible |  | High
| CFG_002 | Verify valid configuration save | Enter valid values in all fields → Click Configure/Save | Configuration saved successfully with confirmation message | Check success popup/message and saved values |  | High
| CFG_003 | Verify blank mandatory field validation | Leave mandatory fields blank → Click Configure | Validation message displayed; no crash | Confirm error message shown and no app closure |  | High |  |  | In the previous version, the app crashed when the fields were left empty.
| CFG_004 | Verify invalid sampling frequency input | Enter letters/special chars/negative value → Click Configure | Proper validation error displayed | Confirm invalid input rejected |  | High
| CFG_005 | Verify minimum sampling frequency acceptance | Enter minimum allowed frequency → Save | Configuration accepted successfully | Confirm value saved without error |  | Medium
| CFG_006 | Verify maximum sampling frequency acceptance | Enter maximum allowed frequency → Save | Configuration accepted successfully | Confirm value saved correctly |  | Medium
| CFG_007 | Verify out-of-range sampling frequency rejection | Enter frequency above supported limit → Save | Error message displayed | Confirm invalid value not accepted |  | High
| CFG_008 | Verify invalid acquisition duration validation | Enter unsupported/negative duration → Save | Validation error displayed | Confirm save blocked |  | High
| CFG_009 | Verify valid acquisition duration | Enter valid duration → Save | Configuration saved successfully | Confirm duration applied correctly |  | Medium
| CFG_010 | Verify invalid interval validation | Enter blank/negative interval → Save | Validation message displayed | Confirm invalid interval rejected |  | High
| CFG_011 | Verify behavior when device disconnects during configure | Start configure → Disconnect device | Disconnect message displayed; no crash | Confirm app stability and error handling |  | High
| CFG_012 | Verify configure without device connection | Open configure without connecting device | App prompts user to connect device | Confirm configure action blocked |  | High
| CFG_013 | Verify multiple configure button clicks handling | Tap Configure button repeatedly | Only one configure request processed | Confirm no duplicate action/request |  | Medium
|  | Verify interruption during configure | Interrupt Bluetooth/network during configure | Proper error handling shown | Confirm no crash or freeze |  | High
| CFG_015 | Verify config persistence after restart | Save config → Close app → Reopen | Saved configuration persists | Confirm saved values retained |  | Medium
| CFG_016 | Verify decimal input rejection | Enter decimal value in integer-only field | Validation error displayed | Confirm invalid input blocked |  | Medium
| CFG_017 | Verify special character validation | Enter @,#,$ in numeric field | Input rejected | Confirm validation works properly |  | Medium
| CFG_018 | Verify app stability during configure | Start configure operation | App should not logout unexpectedly | Confirm stable session |  | High
| CFG_019 | Verify configure timeout handling | Delay device response intentionally | Timeout message displayed | Confirm timeout handled properly |  | Medium
| CFG_020 | Verify configured values reflect in acquisition data | Save config → Start acquisition | Data/graph follows configured settings | Compare waveform/data with configured values |  | High
| CFG_021 | Verify without configured values no reflect in acquisition data | Save config → Start acquisition | No Data/graph follows configured settings | No change in graph |  | High |  |  | Without configuration, the app is plotting the graph once the value is entered.
| Approved By | Mr.Thomas Joseph
| Date | //2026

## DEVICE TESTCASES
TEST CASE REPORT
Project Name : BSH-VAS
Hardware Version | V1.1 |  | Test Summary
Software version | 2.0.2_05 |  | TEST PASS | 0 | #DIV/0! | VAS
Date | 24-05-2026 |  | TEST FAIL | 0 | #DIV/0!
Report Prepared By | Jagatheswaran J |  | TEST BLOCK | 0 | #DIV/0!
Test Module | Test ID | Test Description | Test Procedure | Expected Output | Actual Output | Remarks | Test Status (Pass / Fail / Block)
Power supply | TC_PWR_01 | Verify device power ON with valid 24V | Connect 24V adapter → Turn ON switch | Device powers ON successfully
| TC_PWR_02 | Verify device behavior with low voltage | Supply <24V | Device should not start / show error
| TC_PWR_03 | Verify over-voltage protection | Supply >24V | Device should not get damaged
| TC_PWR_04 | Power interruption test | Disconnect power while running | Device shuts down safely
| TC_PWR_05 | Battery fallback | Remove adapter while ON | Device switches to battery
Accelerometer Inputs | TC_ACC_01 | Verify sensor connection | Connect accelerometer | connect the sensor
| TC_ACC_02 | Validate data acquisition | Provide vibration input | Correct data displayed
| TC_ACC_03 | Channel isolation test | Input only CH1 | CH2 should remain unaffected |  |  |  | Channel 2 wave form is not getting some time.
| TC_ACC_04 | Noise handling | Provide noisy signal | Device filters or logs correctly
| TC_ACC_05 | Sensor disconnection | Remove sensor | Error/alert displayed |  |  |  | No alerts getting
Battery Backup | TC_BAT_01 | Battery charging | Connect power | Battery starts charging
| TC_BAT_02 | Full charge indication | Charge fully | Indicator shows full
| TC_BAT_03 | Backup duration | Run on battery | Meets expected backup time
| TC_BAT_04 | Low battery alert | Drain battery | Warning shown
| TC_BAT_05 | Battery fail case | Disconnect battery | Device should still run on power
Ethernet | TC_ETH_01 | Ethernet connectivity | Connect LAN cable | Link established
| TC_ETH_02 | Data transmission | Send data | Data received correctly
| TC_ETH_03 | Speed validation | Check 10/100 switching | Works correctly
| TC_ETH_04 | Disconnect handling | Remove cable | Device detects disconnection
| TC_ETH_05 | IP configuration | Set static(default) | works properly
USB Host | TC_USB_01 | USB detection | Insert USB device | Device detects USB
| TC_USB_02 | File read/write | Access USB files | Data read/write success
| TC_USB_03 | Unsupported device | Insert unsupported USB | Error shown
| TC_USB_04 | Hot plug test | Insert/remove USB | No crash
| TC_USB_05 | Power supply check | Connect USB device | Proper power given
Acquisition Start Button | TC_BTN_01 | Start acquisition | Press button(external trigger) | Acquisition starts |  |  |  | App is in GUI trigger mode, but when External Trigger is turned OFF, the device does not post data .
| TC_BTN_02 | Multiple press handling | Press repeatedly | No crash
| TC_BTN_03 | Debounce test | Press quickly | Works correctly
Reset Button | TC_RST_001 | Verify device resets when reset switch is pressed | Device is powered ON and functioning normally | Device restarts
| TC_RST_002 | Verify behavior on long press reset | Press and hold reset switch for 5–10 seconds | No abnormal behavior
| TC_RST_003 | Reset During Operation | Reset while device performing active task | Operation stops immediately
| TC_RST_004 | Reset During Boot | Press reset during boot-up | Device restarts boot process cleanly
| TC_RST_005 | Power + Reset Combination | Reset behavior with power cycling | Device enters recovery
| TC_RST_006 | Multiple Rapid Press | Press reset switch rapidly multiple times | No unexpected behavior
Marked yellow not implemented or tested

## TPP
Stage | Test | Description | Expected result | Observed Result | Date | HW | ES | AS | OS | System Config
Configurable sampling frequency | [9 empty] | RAM ….
1 | 1 | Conduct as per "ConfigFs" worksheet
Impact of acquisition duration
2 | 1 | Conduct as per "Acq.Dur" worksheet
Accuracy of Acquisition interval and Duration
3 | 1 | Conduct as per "AcqIntDur" worksheet
Accuracy of GUI Parameters
4 | 1 | Conduct as per APP Testcases worksheet
Accuracy of GUI Parameters
5 | 1 | Conduct as per Device TestCases

## ConfigFs
Test Objective : Configurable sampling frequency and Time series & Spectrogram plots
This test to be conducted without device power cycle/ device restart / application restart. If done , specify in remarks
Sl No | Frequency Generator Input Freqency (Hz) | Frequency Generator Input Amplitude (V) | Frequency Generator Input waveform | Sampling Frequency (Hz) | Interval (Sec) | Duration (Sec) | Cycle duration in Time Series (milisec) | Spectrogram Main Peak (Hz) | Spectrogram other peaks (Hz) | Spectrogram Main Peak (Amplitude) | Remarks | Status | Observations
1 | 0.1 | 4 | Sine | 1000 | 20 | 12
2 | 1 | 4 | Sine | 1000 | 20 | 2
3 | 5 | 4 | Sine | 2000 | 10 | 1
4 | 100 | 4 | Sine | 2000 | 10 | 1
5 | 1000 | 4 | Sine | 10000 | 5 | 0.1
6 | 3000 | 4 | Sine | 10000 | 5 | 0.1
7 | 3000 | 4 | Sine | 15000 | 5 | 0.1
8 | 5000 | 4 | Sine | 22000 | 5 | 0.1
9 | 5000 | 4 | Sine | 40000 | 5 | 0.1
10 | 5000 | 4 | Sine | 80000 | 5 | 0.1
11 | 10000 | 4 | Sine | 22000 | 5 | 0.1
12 | 10000 | 4 | Sine | 40000 | 5 | 0.1
13 | 10000 | 4 | Sine | 80000 | 5 | 0.1
14 | 18000 | 4 | Sine | 40000 | 5 | 0.1
15 | 18000 | 4 | Sine | 80000 | 5 | 0.1
16 | 25000 | 4 | Sine | 80000 | 5 | 0.1
17 | 0.1 | 4 | Square | 1000 | 20 | 12
18 | 1 | 4 | Square | 1000 | 20 | 2
19 | 5 | 4 | Square | 2000 | 10 | 1
20 | 100 | 4 | Square | 2000 | 10 | 1
21 | 1000 | 4 | Square | 10000 | 5 | 0.1
22 | 3000 | 4 | Square | 10000 | 5 | 0.1
23 | 3000 | 4 | Square | 15000 | 5 | 0.1
24 | 5000 | 4 | Square | 22000 | 5 | 0.1
25 | 5000 | 4 | Square | 40000 | 5 | 0.1
26 | 5000 | 4 | Square | 80000 | 5 | 0.1
27 | 10000 | 4 | Square | 22000 | 5 | 0.1
28 | 10000 | 4 | Square | 40000 | 5 | 0.1
29 | 10000 | 4 | Square | 80000 | 5 | 0.1
30 | 18000 | 4 | Square | 40000 | 5 | 0.1
31 | 18000 | 4 | Square | 80000 | 5 | 0.1
32 | 25000 | 4 | Square | 80000 | 5 | 0.1

## Aqtn Duration
Test Objective : Configurable Acquisition Duration
This test to be conducted without device power cycle/ device restart / application restart. If done , specify in remarks
Sl No | Frequency Generator Input Freqency (Hz) | Frequency Generator Input Amplitude (V) | Frequency Generator Input waveform | Sampling Frequency (Hz) | Duration (Sec) | Interval (Sec) | No of Wave in 1ms | Remarks | Status
1 | 1000 | 4 | Sine | 22000 | 1 | 5 | 1 | SINE WAVE
2 | 3000 | 4 | Sine | 22000 | 1 | 5 | 3 | Distorted sine wave
3 | 5000 | 4 | Sine | 22000 | 1 | 5 | 5 | Distorted sine wave
4 | 10000 | 4 | Sine | 22000 | 1 | 5 | 10 | Distorted sine wave
5 | 1000 | 4 | Square | 22000 | 1 | 5 | 1 | Distorted square wave
6 | 3000 | 4 | Square | 22000 | 1 | 5 | 3 | Distorted square wave
7 | 5000 | 4 | Square | 22000 | 1 | 5 | 5 | Distorted square wave
8 | 10000 | 4 | Square | 22000 | 1 | 5 | 10 | Distorted square wave

## Aqt Int_Duration
Test Objective : Configurable Acquisition Interval
This test to be conducted without device power cycle/ device restart / application restart. If done , specify in remarks
Sl No | Frequency Generator Input Freqency (Hz) | Frequency Generator Input Amplitude (V) | Frequency Generator Input waveform | Sampling Frequency (Hz) | Duration (Sec) | Interval (Sec) | Log File Count for 1min | Remarks | Status
1 | 1000 | 4 | Sine | 22000 | 1 | 5 | 7
2 | 3000 | 4 | Sine | 22000 | 1 | 10 | 4
3 | 5000 | 4 | Sine | 22000 | 1 | 5 | 4
4 | 10000 | 4 | Sine | 22000 | 1 | 10 | 4
5 | 1000 | 4 | Square | 22000 | 1 | 5 | 12
6 | 3000 | 4 | Square | 22000 | 1 | 10 | 4
7 | 5000 | 4 | Square | 22000 | 1 | 5 | 5
8 | 10000 | 4 | Square | 22000 | 1 | 10 | 4
Test Objective : Configurable Acquisition Duration
This test to be conducted without device power cycle/ device restart / application restart. If done , specify in remarks
|  | Sl No | Frequency Generator Input Freqency (Hz) | Frequency Generator Input Amplitude (V) | Frequency Generator Input waveform | Sampling Frequency (Hz) | Duration (Sec) | Interval (Sec) | No of Wave in 1ms | Status
|  | 1 | 1000 | 4 | Sine | 22000 | 1 | 5
|  | 2 | 3000 | 4 | Sine | 22000 | 1 | 5
|  | 3 | 5000 | 4 | Sine | 22000 | 1 | 5
|  | 4 | 10000 | 4 | Sine | 22000 | 1 | 5
|  | 5 | 1000 | 4 | Square | 22000 | 1 | 5
|  | 6 | 3000 | 4 | Square | 22000 | 1 | 5
|  | 7 | 5000 | 4 | Square | 22000 | 1 | 5
|  | 8 | 10000 | 4 | Square | 22000 | 1 | 5